## Slide 1
CalxPy: a software for the calibration of geophysical systems
against a reference
Benoît Doury1, Ichrak Ketata2
1 CTBTO Preparatory Commission, Vienna, Austria
2 Zühlke Engineering GmbH
Poster P3.1-243

## Slide 2
P3.1-243
ABSTRACT
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
The IMS Operational Manuals for waveform stations require that IMS stations be calibrated regularly. Since 2012, the PTS had relied mostly on electrical calibration to meet that requirement.
However electrical calibration comes with some challenges (no traceability, integration and sustainment issues, high operating costs…).
A part of the geophysical community, including Station Operators, has started performing regular calibrations by comparison against a co-located reference. This method allows a more systematic and centralized approach to calibration. Over the past few years it has been gradually more used at IMS stations, particularly infrasound. In this context, the PTS is developing tools to support this alternative approach.
Here we present CalxPy, a web-application developed at the PTS for the calibration of geophysical systems by comparison. With CalxPy, one can calculate, store and display the response of a system for a given period, or track the evolution of the response against time or environmental variables. CalxPy also allows the refinement and evaluation of the measured response against a baseline, and the reporting of IMS2.0 calibration results.
CalxPy supports the Initial calibration and On-site yearly calibration processes, as well as Data Quality Control.
CalxPy can be deployed in the IDC pipeline and in NDC-in-a-box.

## Slide 3
P3.1-243
INTRODUCTION
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
CTBT/WGB/TL-11,17/17/Rev.5
Extensive testing of a new device model against IMS and manufacturer model specifications
Type Approval
Acceptance Testing
Initial Calibration
On-site Calibration
Testing of an individual device against manufacturer model specifications
  Testing (full frequency response + self-noise) of a new system in operational conditions against IMS specifications
  Yearly measurements of full frequency response of operational system and comparison against initial calibration results
when equipment change
S/N S000x
S/N D000x
S/N W000x

## Slide 4
P3.1-243
INTRODUCTION
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
2014: passive calibration method using side-by-side comparison validated for infrasound technology1
2015: development and tuning of a prototype in Matlab2
2016: prototype ported to Python scripts and streamlined
validation against initial Matlab prototype (results accuracy)
performance optimized
2017-2018: integration into IDC environment
Integration into internal technology web portal for results representation
Automate runs (monthly and every other day)
1 Gabrielson, 2014
2 Charbit, 2015
2019-2021: Python scripts integrated into a full software solution: CalxPy
Closer match with IMS processes
Adapt software to new non-functional requirements
Develop an interactive frontend to visualize raw data
History

## Slide 5
P3.1-243
METHODS
Raw processing overview
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
H1
C1
IDC
Apply filter k
1
2
3
4
5
6
[0.01-0.06 Hz]
[0.05-0.15 Hz]
[0.14-0.35 Hz]
[0.32-0.7 Hz]
[0.6-1.5 Hz]
[1.5-6Hz]
1000 s
250 s
150 s
30 s
select time intervals ik
xH,1,1(t)
xC,1,1(t)
xH,1,2(t)
xC,1,2(t)
xH,1,3(t)
xC,1,3(t)
xH,1,4(t)
xC,1,4(t)
xH,1,5(t)
xC,1,5(t)
50% overlap
700 s
xH,1,6(t)
xC,1,6(t)
20 s
Response
equation
MSCHC,k,i(f)
xH,k,i(t)
xC,k,i(t)
CPSD
PSD
PSD
ᴕHH,k,i(f)
ᴕCC,k,i(f)
ᴕHC,k,i(f)
Coherence
equation
GH,k,i(f)
Loop on (k,i)
Frequency Response
Coherence
Raw results

## Slide 6
P3.1-243
METHODS
Composite Processing overview
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
ᴕHH,k,i(f)
ᴕHC,k,i(f)
ᴕCC,k,i(f)
Composite results
Raw
results
2
1
Masking Criteria
I37H0 - 2021
Amplitude
Phase
Amplitude (Pa/cnt)
@0.25Hz

## Slide 7
P3.1-243
METHODS
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
Functional Requirements
Initial and Scheduled Calibration
CALIB / NCALIB
Full frequency response
Evaluation
IMS Operational Manuals
IMS2.0 Formats and Protocol
Quality Assurance documents
IN_SPEC YES / NO
Quality Control
Raw Results exploration
Continuous tracking
Detect configuration changes
Identify trends, seasonal patterns
Sensitivity to environment
Alerts
IMS2.0

## Slide 8
P3.1-243
METHODS
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
Non-Functional Requirements
Scalable
Can run in distributed IDC environment (300+ site pairs), but also on a field laptop (NDC in a box)
Integrated
Provides a standard interface to other IDC software (for the scheduling, evaluation and publication of results)
Versatile
Can be used with other waveform technologies Support CSS and MiniSEED
Smooth
Exploration and Interactive visualization of Raw and Composite results is smooth for the end user
Reproducible
Processing parameters used to obtain results are enough documented to allow reproducibility and re-use on new data sets
Tested and Testable
Core of the algorithm is fully tested and testable
Trustworthy
Guaranties integrity of published results (authentication, safe transactions with VDMS, …)

## Slide 9
P3.1-243
Get latest from Result Store
Task Progress / Computed Values
Get Task Status by ID
Frontend
Backend Server
RabbitMQ
Celery Worker
Result Store
Requirements
The user can start multiple computations in parallel
The user can continue navigating during computation
Start Task
Task ID
Add to queue
Queue
Get Task to be executed
Save Status
Save Results (Authorization)
PostgreSQL DB
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
METHODS
Software engineering
Example of challenge: Computing composite responses “on-the-fly”
Solution: Asynchronous Computation
Challenges
The computation includes big amounts of data (up to 1.6 millions records for 1 year of data)
The computation can take several minutes
A task ID is generated and sent back to the frontend
The user starts new computation from the web interface
The task is added to the queue of the message broker RabbitMQ
The Celery worker consumes this task and starts the computation
The status of computation is saved to the file system
The frontend application checks the status of execution periodically and updates the plot

## Slide 10
P3.1-243
RESULTS
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
Scheduled Calibration
Statistics
Historical statistics
Station / Channel Views
Process, Results and Publication statistics
Select a year for review

## Slide 11
Mark for expert review?
Publish result (IMS2.0 or JSON)
P3.1-243
RESULTS
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
Scheduled Calibration
Station view
I37H0/BDF
I37H1/BDF
I37H2/BDF
I37H3/BDF
I37H4/BDF
I37H5/BDF
I37H7/BDF
Quick overview of station calibration
I37H8/BDF
I37H9/BDF
I37H6/BDF
1.
2.
3.
4.
5.
Expand (see next slide)
Auto-evaluation
Reviewer makes his evaluation
6.

## Slide 12
Mark for expert review
Publish result (IMS2.0 or JSON)
P3.1-243
RESULTS
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
Scheduled Calibration
Expand a site
Display points out of +/-5% tolerance
3.
2.
5.
6.
Visualize missing data
Change dates selection
Recompute
4.

## Slide 13
P3.1-243
CONCLUSIONS
CalxPy: a software for the calibration of geophysical systems against a reference
Benoît Doury, IMS division, benoit.doury@ctbto.org
Ichrak Ketata, Zuehlke Engineering GmbH, ichrak.ketata@zuehlke.com
The geophysical community, including Station Operators, is performing regular calibrations by comparison against a co-located reference
The PTS is developing the CalxPy software to support this approach
CalxPy closely supports PTS processes for Quality Assurance (Device Acceptance, Initial Calibration, Scheduled Calibration) and Quality Control
CalxPy is a web application designed to allow for a smooth and interactive user experience even when running against very large datasets (300+ site pairs over 10 years)
Results in CalxPy can be exported in a way (JSON) that allows reproducibility of results
CalxPy can be deployed in the IDC pipeline (distributed) and in NDC-in-a-box (standalone)