## Slide 1
A coherent gas-combustion infrasound source
Chad Smith and Tom Gabrielson
Email: chad.smith@psu.edu
This work was supported by the Nuclear Arms Control Technology (NACT) Program at the Defense Threat Reduction Agency (DTRA).
Cleared for Release
CTBT Science and Technology 2021
Poster No: P3.1-713

## Slide 2
Primary goal: Development of a practical coherent (non-explosive, periodic, with controlled duration) infrasound source to aid in IMS verification, calibration, and coherence studies.
Frequency coverage from 0.1 Hz to several Hz
Suitable for coherent processing (helping to increase SNR over wind noise)
Usable signal level at ranges greater than 1 km
Transportable
Reasonably cost efficient
Design goals:
Smith, C. M., & Gabrielson, T. B. (2020). Farfield coherent infrasound generation using an air-propane burner. The Journal of the Acoustical Society of America, 148(5), 3181-3194.
Smith, C. M., & Gabrielson, T. B. (2018). Measurements and modelling of a gas-combustion infrasound source. In Proceedings of Meetings on Acoustics 176ASA (Vol. 35, No. 1, p. 045004). Acoustical Society of America.
Smith, C. M., & Gabrielson, T. B. (2018, April). Scaling of a gas-combustion infrasound source. The Journal of the Acoustical Society of America, 143(3), 1808-1808.
Smith, C. M., & Gabrielson, T. B. (2015). Theoretical model and experimental validation of a gas-combustion infrasound source. The Journal of the Acoustical Society of America, 137(4), 2407-2407.
Further information:

## Slide 3
volume acceleration
Any practical (transportable) infrasound source will be much smaller wave length in the infrasound regime. However, simple sources are commonly known to be highly inefficient. The pressure from a simple source assumed to be baffled by the earth is described as,
This relationship shows how the time domain pressure for a simple source is proportional to the volume or mass acceleration. Assuming a time harmonic source, examining the amplitude of this relationship in the frequency domain,
it is obvious that (due to a frequency squared relationship) very large amounts of air must be displaced to create sub-hertz signals as useful levels.
Practical infrasound sources are inefficient

## Slide 4
Due to the fundamental inefficiency of simple sources, the authors have proposed an unconventional but practical method of coherent infrasound excitation— using gas combustion to periodically displace large volumes of air within the open atmosphere.
While an infrasound source will be inherently inefficient, this source method can increase the useful range of a sub-hertz infrasound source compared with previous work.
Available energy density
Custom Built 10 MW Infrasound Source
Periodic displacement of gaseous combustion constituents not only allows far greater displacement than a conventional transducer, but also has four orders of magnitude greater energy density than releasing compressed air at common pressure levels.
Potential energy density in air at
150 PSI: 2 MJ/m3
Combustion energy density in propane: 29,000 MJ/m3

## Slide 5
cp = specific heat at constant pressure for CM
M = molecular weight
P0 = atmospheric pressure
= volume velocity
R = universal gas constant
= effective burner heat power
CM
Thermodynamic model
An idealized thermodynamic model has been developed to estimate the volume velocity output of a cycled gas-burner. This model defines a thermodynamic control-mass (CM) encompassing the burner’s gaseous volume injection of highly heated combustion constituents. Assuming stoichiometric combustion, the volume velocity output of the source is directly proportional to the effective heat power of the burner. This relationship displays the scalable nature of this method, e.g., two burners placed near one another would output twice the volume velocity of a single burner. Please see reference at the bottom of this slide for model development.
Smith, C. M., & Gabrielson, T. B. (2020). Farfield coherent infrasound generation using an air-propane burner. The Journal of the Acoustical Society of America, 148(5), 3181-3194.

## Slide 6
The source was designed to maximize heat power output while ensuring safe and automated control. The current prototype is compose of dual liquid propane burners. Liquid propane is periodically injected into the vaporization coils using digitally controlled solenoid valves.
50 to 1000 meters from the source. The source was operated for five minutes synchronously and the entire five minutes was processed using a single DFT.
Source design and field testing
Custom Built 10 MW Infrasound Source
A real-time-capable embedded controller in conjunction with solenoid valve control circuitry allows each burner to be tightly and synchronously controlled. This system also allows the system to be controlled remotely via an Ethernet/SSH connection.
Field testing was done by placing calibrated GRAS infrasound microphones throughout the test site (slide eight) with 3 or 6 meter wind-noise reduction systems. Microphones were placed from
vaporization coils
vapor fuel ports

## Slide 7
Source operation requires a separate 43 lb liquid propane tank for each burner. These tanks are similar to those commonly used to power forklifts, however, these tanks have been retrofitted with high flow rate valves for liquid propane extraction, and vapor injection valves. The vapor injection valves are used to inject nitrogen
Source operation and measurements
Field-measurement Configuration
into the top of the tank. While nitrogen is inert and does not combust, this allows careful control of propane fuel delivery pressure. This control decreases the impact of latent heat of vaporization chilling the fuel tanks, increases source output, and reduces variation in source output signal level.
A standard vapor propane tank is used to power each burner’s pilot light. In early tests, a ladder and force gauge where used measure
fuel usage, while a heat flux sensor was used to monitor thermal radiation. The ladder, force gauge, and heat flux sensor are not required for source operation.

## Slide 8
200 m
receivers at 100, 200, 400, and 1000 meters
Russell E. Larson
Agricultural Research Center
Penn State University
Source/receiver locations
3-6 m PVC Wind-noise Reduction Systems

## Slide 9
purpose-built
double-burner
heat flux sensor
force gauge
(fuel usage monitor)
nitrogen injection
system
vapor propane for pilots
liquid propane tanks
storage case
Measurement
Configuration

## Slide 10
fundamental and many harmonics still above noise floor at 1km
Signal-to-noise vs. Range (0.25 Hz fundamental)

## Slide 11
5.25 dB
Greatest Source Output (2019 vs. 2020)
Increasing useful source range
In the 2019 test season, the thermodynamic engineering model and source method were validated. It was experimentally shown that this source method can produce sub-hertz infrasound that propagates into the farfield.
The 2020 test season focused on improving acoustic output through adjusting vapor fuel port dimensions, fuel delivery pressure, and the duty cycle of fuel modulation. While 2020 testing used a single burner and ranges of ≤400 m, acoustic output is believed to have been increased 6-12 dB at the fundamental.
(100 m range)

## Slide 12
Further refinement of source operations and optimization of useful range
Characterization of frequency band capability for the fundamental (testing thus far has been at 0.25 Hz)
On-site testing at an IMS or similar monitoring station
Future work:
A large gas-combustion infrasound source prototype was designed, built, and experimentally verified to create infrasound that propagates into the farfield
This source is easily transported and suitable for phase-coherent processing of signals to help overcome wind noise
Usable signal level was recorded at ranges greater than 1 km in low wind conditions in 2019. More recent work on this topic is believed to have increased signal level 6-12 dB compared with 2019 measurements.
Results: