## Slide 1: Operation and maintenance of Japanese RN stations under the COVID-19 related
Yutaka Tomita, Akiko Furuno, Yoichi Yamamoto
CTBT and Nuclear Materials Transport Support Section
Integrated Support Center for Nuclear Nonproliferation and Nuclear Security
Japan Atomic Energy Agency

## Slide 2: Radionuclide Facilities in Japan
【Radionuclide Stations】
・Okinawa (RN37)
・Takasaki (RN38)
【Radionuclide Laboratory】
・Tokai (RL11)
【National Data Center-2】
・Tokai
Station operator & station
manager stay in Tokai
Local operators perform
daily checks and maintenance
works on site
Tokai Radionuclide Laboratory（RL11）
Takasaki Station
（RN38）
・Particle(JPP38)
・Noble gas(JPX38)
Okinawa Station
（RN37）
・ Particle(JPP37)
NDC-2
(for Radionuclide)

## Slide 3: Impact of “State of Emergency” for the station
First wave
National state of
emergency
(4/16-5/25)
Second wave
Third wave
Fourth wave
Number of positive with PCR test in Japan
Japanese government asked for a minimum of reduction in people-to-people contact and performed immigration restrictions of visitors from foreign countries.
State of emergency
for metropolitan area
(1/8-2/21)
・Reduce local operators and JAEA staffs to less than half
・Self-restraint of business trip for maintenance work
・Postponement of maintenance performed by manufacturer
State of emergency
for some areas
(4/23-6/20)
Chart: 
JAEA's response to the government's request

## Slide 4
Trouble shooting for RN37 and RN38
RN37 Okinawa (JPP37)
・Failed a preamplifier of detector
・Failed a System Instrumentation Unit
RN38 Takasaki
1)JPP38
・ Failed a preamplifier of detector
・RASA system down due to barcode
misreading
CTBTO
Local Operator
Manufacturer
・General Dynamics :USA
(JPP37, JPP38)
・Scienta SAUNA systems :Sweden
(JPX38)
Station Manager
Station Operator
Takasaki,
Okinawa
Tokai
Report
Report
Trouble
Shooting
Trouble
Shooting
Technical Advice
2)JPX38
・Failed sampling pumps
Flowchart for trouble shooting
Troubles of RN37 and 38 during COVID-19 pandemic
Information
Sharing

## Slide 5: Example-1 Trouble caused by postponement of annual maintenance of SAUNA
Normally, an annual maintenance for the SAUNA is carried out as followings with an engineer of the manufacture (Scienta SAUNA system)
Main annual maintenance works
・General system check-up and evaluation of data
・ Replacement of part and equipment
sampling pumps, air inlet filter, check
valves and gas line filters, etc
・ Leak check of the system
・ Recalibrate of the Gas chromatography
・ Check of detector performance
Sampling pressure drop
Sampling pumps were failed due to postpone of
the annual maintenance
※ SAUNA : Swedish Automatic Unit for Noble Gas Acquisition
※

## Slide 6: Example-1 Replacement of pumps by local operators
Detailed instruction sent from CTBTO
Replacement of the sampling pumps

## Slide 7: Example-2 Trouble of a detector for RASA
Noise
・Many noises appeared in the spectrum of RASA
・The cause was corrosion of the preamplifier
in the detector
※RASA：Radionuclide Aerosol Sampler/Analyzer
※

## Slide 8
RN38 Takasaki
Example-2
Remote support for maintenance work using web conferencing software
Tokai Office
Web camera
Monitor
Connecting with
Web Conferencing
Software
Advantages of using the web conferencing software
・Instructions of maintenance work can be given to local operators remotely
・Save time and travel allowance
・Share the maintenance experience with other staffs
・Recording can be used as training material for current and future operators

## Slide 9: Summary
Equipment failures that occurred during COVID-19 pandemic were addressed through close communication with the stations, CTBTO, and manufacturers.
Instructions sent from CTBTO are very detailed and helpful. Station Operators hope that more and more such instructions will be developed.
It is very effective to use web conferencing software for support of the maintenance work done by local operators.